POCKETLAB

# Bipolar AC Output Stage

Breadboard Validation and Characterization Report

| Project | PocketLab Programmable Function Generator |
| --- | --- |
| Test date | July 22, 2026 |
| Prepared for | Mark Anderson |
| Status | PASS — provisional high-impedance AC output validated |

Executive result A 47 µF coupling capacitor, 100 Ω protection resistor, and permanent 100 kΩ ground-reference resistor converted the existing 2.34 V-biased output into a clean bipolar waveform centered within a few millivolts of 0 V. Amplitude tracking was within 0.5% from 10 Hz through 100 kHz under high-impedance loading. Loads of 10 kΩ or greater produced about 1% or less additional attenuation.

# 1. Objective

Validate a passive output branch that removes the function generator’s DC bias while preserving the AC waveform. The stage was evaluated for waveform centering, amplitude transfer, frequency response, load response, and power-switching behavior.

# 2. Circuit under test

| Connection | Component / node | Purpose |
| --- | --- | --- |
| DC OUT → capacitor + | 47 µF, 10 V electrolytic | Blocks steady DC after charging; passes the changing waveform |
| Capacitor − → AC OUT | 100 Ω series resistor | Limits fault and capacitive-load current; creates predictable load attenuation |
| AC OUT → ground | 100 kΩ permanent resistor | References output average to 0 V and discharges the capacitor |
| Measurement | Analog Discovery 3, 1 MΩ input | High-impedance observation of DC OUT and AC OUT |

Capacitor polarity was positive toward the OP484 output and negative toward AC OUT. The existing DC output remained available and unchanged.

# 3. Operating principle

At steady state, the capacitor charges to approximately the 2.34 V DC bias at DC OUT. The constant component is blocked, while voltage changes are coupled through. The 100 kΩ resistor establishes ground as the new average level.

vDC(t) ≈ 2.343 V + 0.496 sin(2πft) V

vAC(t) ≈ 0.496 sin(2πft) V

The permanent 100 kΩ resistor in parallel with the 1 MΩ instrument input gives about 90.9 kΩ. With 47 µF, the calculated high-pass corner is approximately 0.037 Hz—far below the tested 10 Hz minimum.

Important limitation This stage produces negative waveform excursions but not a negative DC rail. It supports an AC waveform centered at 0 V; it cannot provide steady negative offset or an adjustable center below ground.

# 4. Safety and initial powered checks

Continuity confirmed from OP484 pin 1 / existing output to the capacitor positive lead.

Capacitor negative lead to AC OUT measured approximately 99.1 Ω.

AC OUT to circuit ground measured approximately 99.8 kΩ.

AC OUT to +5 V measured about 100 kΩ and rising, consistent with capacitor charging rather than a short.

+5 V to ground showed no short; measurement grounds used common circuit ground.

Initial powered checks were normal: +5 V rail = 4.71 V; supply current = 0.06–0.07 A; DC OUT = 2.345 V when active; AC OUT decayed toward 0 V. With output logically off, DC OUT measured 10.6 mV and AC OUT measured 0.0 mV.

# 5. Initial 1 kHz validation

| Metric | AC OUT result | Interpretation |
| --- | --- | --- |
| Maximum | +0.50396 V | Near expected +0.5 V peak |
| Minimum | −0.50196 V | Near expected −0.5 V peak |
| Peak-to-peak | 1.0059 V | 0.59% above 1.00 V command |
| Average | +1.818 mV | Essentially centered at ground |
| Frequency | 999.05 Hz | Matches 1 kHz command |

A two-channel capture showed DC OUT at approximately 0.992 Vpp centered at 2.343 V and AC OUT at approximately 0.994 Vpp centered near 0 V. The stage removed the bias while preserving waveform amplitude and phase.

# 6. No-added-load frequency sweep

| Set f | AC Vpp | DC Vpp | AC/DC | AC average | Phase |
| --- | --- | --- | --- | --- | --- |
| 10 Hz | 0.9473 | 0.9505 | 99.67% | +1.72 mV | 0.254° |
| 100 Hz | 0.9560 | 0.9606 | 99.52% | +1.92 mV | 0.245° |
| 1 kHz | 0.9936 | 0.9952 | 99.84% | +1.97 mV | 0.205° |
| 10 kHz | 0.9966 | 0.9982 | 99.84% | +2.00 mV | 0.311° |
| 50 kHz | 0.9874 | 0.9888 | 99.86% | −0.64 mV | 0.473° |
| 100 kHz | 0.9487 | 0.9504 | 99.83% | +1.72 mV | 0.552° |
| 200 kHz | 0.8427 | 0.8310 | 101.41%* | +2.60 mV | 0.794° |

*At 200 kHz the waveform was visibly distorted; automatic Vpp is sensitive to peaks and noise, so the apparent ratio above 100% is not passive gain.

## Frequency-sweep findings

AC OUT remained within approximately −0.6 mV to +2.6 mV of ground.

From 10 Hz through 100 kHz, AC OUT differed from DC OUT by less than 0.5%.

Phase difference stayed below 0.8° at every tested point.

No visible low-frequency droop occurred at 10 Hz.

Amplitude reduction and distortion near 200 kHz appeared on both outputs and therefore originate primarily in the upstream DDS/amplifier path.

# 7. 1 kHz load sweep

| Added load | Effective load* | AC Vpp | DC Vpp | AC/DC | AC average |
| --- | --- | --- | --- | --- | --- |
| None | 90.9 kΩ | 0.9940 | 0.9955 | 99.85% | +11.28 mV |
| 100 kΩ | 47.6 kΩ | 0.9893 | 0.9917 | 99.76% | +4.11 mV |
| 10 kΩ | 9.01 kΩ | 0.9831 | 0.9925 | 99.06% | +2.16 mV |
| 1 kΩ | 988 Ω | 0.8932 | 0.9858 | 90.61% | +0.84 mV |

*Includes the permanent 100 kΩ resistor and the Analog Discovery’s 1 MΩ input.

Loss closely matches the divider formed by the intentional 100 Ω resistor and effective load. For the 1 kΩ test, predicted transfer was 90.8% and measured transfer was 90.6%. Estimated peak output current was approximately 0.45 mA.

## Load-sweep findings

Loads of 10 kΩ or greater preserved amplitude to about 1% or better.

The 1 kΩ load caused approximately 9.4% attenuation, consistent with the 100 Ω protection resistor.

No clipping, oscillation, or major distortion was visible under any tested load.

DC OUT remained stable; the loads did not significantly disturb the OP484 stage.

The output is not intended for 50 Ω loads.

# 8. Power-switching transient

A separate bench-supply switching capture showed a positive transient at both outputs. The largest recorded excursions were approximately +2.80 V at AC OUT and +2.88 V at DC OUT, with no significant negative excursion. These remained inside the approximately 0–4.71 V internal supply range.

The coupling capacitor converts a rapid internal DC-level change into a pulse at AC OUT. Narrow spikes and the intermediate plateau may also include supply switching or switch bounce. The capture does not prove that extremely short spikes are absent, but it showed no immediately dangerous excursion for the generator or Analog Discovery.

External-circuit caution The transient is acceptable for breadboard characterization but could disturb a sensitive device connected to AC OUT. Keep the external load disconnected during power transitions, or set OUTPUT OFF and isolate the device before cycling bench power.

# 9. Conclusions and provisional rating

| Parameter | Provisional rating / conclusion |
| --- | --- |
| Waveforms | Sine validated; triangle characterization pending |
| Recommended frequency | 10 Hz–100 kHz |
| Extended operation | 200 kHz with reduced amplitude and visible upstream distortion |
| Amplitude | 1.00 Vpp characterized; higher-amplitude sweep pending |
| Output center | Approximately 0 V; within a few millivolts in the sweep |
| Recommended load | 10 kΩ or greater |
| 1 kΩ operation | Usable with approximately 9% attenuation at 1 kHz |
| 50 Ω drive | Not supported |
| DC capability | AC only; no sustained negative DC or adjustable negative center |

Overall result: PASS. The passive bipolar AC-output stage removes the midrail bias, centers the waveform near ground, and preserves signal integrity over the intended high-impedance range.

# 10. Recommended next steps

Capture startup and shutdown separately at 20–100 ms/div with OUTPUT OFF before cycling power.

Run triangle checks at 10 Hz, 1 kHz, 50 kHz, and 200 kHz.

Test 2.0, 4.0, and 4.15 Vpp into high impedance; verify symmetry and clipping margin.

Add an AC/DC selector, first manual, then evaluate a relay or suitable analog switch with startup isolation.

Implement safe switching: mute, establish midrail, switch coupling, settle, restore amplitude, and unmute.

Add COUPLING=AC/DC to firmware/BLE and update the app so AC mode displays 0 V center and disables adjustable offset.

# 11. Safe breadboard operation

Before shutdown: set OUTPUT OFF, verify AC OUT is near 0 V, isolate the external device, then switch off bench power.

At startup: leave output disconnected, power the generator, wait for safe initialization, then connect the load and enable output.

Do not connect laptop USB and external 5 V simultaneously until the ESP32 board’s dual-power behavior is verified.

Use common ground and avoid wiring changes while powered or attached to the laptop.