PocketLab Safe-State
Regression Test

Final Test Results

| Project | PocketLab BLE-Controlled Function Generator |
| --- | --- |
| Firmware | PocketLab-FG.safe-idle.ino, version 0.3.2 |
| Controller | ESP32 NodeMCU-32S |
| Signal source | AD9833 DDS module |
| Test instrument | Digilent Analog Discovery / WaveForms |
| Date and time | July 13, 2026, 2:19 PM Pacific Daylight Time (PDT) |

OVERALL RESULT: PASS

All eight regression tests passed. The hardware boots safely, STOP and BLE disconnect force the AD9833 into safe-off mode, reconnection does not restart the output, and invalid DC output requests are rejected.

# 1. Test Objective and Configuration

The purpose of this regression test was to verify that the PocketLab function generator remains in a safe physical and logical state during power-up, normal operation, STOP, BLE disconnect, automatic reconnect, and invalid output requests.

## Safe idle state

WAVE=DC
FREQ=0
AMP=0.00
OFFSET=0.00
OUTPUT=OFF

## Wiring

| AD9833 | ESP32 NodeMCU-32S |
| --- | --- |
| VCC | 3.3 V |
| GND | GND |
| SDATA | GPIO 23 |
| SCLK | GPIO 18 |
| FSYNC | GPIO 5 |
| OUT | Analog Discovery scope input |

The ESP32, AD9833, and Analog Discovery shared a common ground.

# 2. Measurement Summary

| Test condition | Maximum | Minimum | Average | Peak-to-peak | Frequency |
| --- | --- | --- | --- | --- | --- |
| Powered-down baseline | 24.283 mV | -5.517 mV | 7.194 mV | 29.800 mV | No valid signal |
| 1. Boot safe state | 44.150 mV | 17.059 mV | 30.828 mV | 27.091 mV | No valid signal |
| 2. Connected idle | 10.116 mV | -3.450 mV | 3.361 mV | 13.566 mV | No valid signal |
| 3. Settings applied, stopped | 13.902 mV | -6.921 mV | 3.655 mV | 20.822 mV | No valid signal |
| 4. 1 kHz sine running | 0.64378 V | 44.346 mV | 343.71 mV | 0.59943 V | 999.97 Hz |
| 5. STOP safe-off | 5.0679 mV | 2.4651 mV | 3.7271 mV | 2.6028 mV | No valid signal |
| 6a. Active before disconnect | 0.64370 V | 44.268 mV | 343.79 mV | 0.59943 V | 999.94 Hz |
| 6b. After BLE disconnect | 5.1468 mV | 2.1496 mV | 3.7469 mV | 2.9972 mV | No valid signal |
| 7. Reconnected, output off | 5.1468 mV | 2.4651 mV | 3.8713 mV | 2.6817 mV | No valid signal |
| 8. Invalid DC RUN | 5.3045 mV | 2.4651 mV | 3.8443 mV | 2.8394 mV | No valid signal |

Note: WaveForms displayed high-frequency values during safe-off measurements. Because these captures were only a few millivolts peak-to-peak and contained no coherent generated waveform, those readings were treated as noise-derived estimates rather than valid output frequencies.

# 3. Detailed Test Results

## Test 1 - Power-up safe state

The ESP32 was power-cycled without connecting the app. The firmware issued the AD9833 safe-off control word and began BLE advertising.

[AD9833] SPI write 0x21C0
[AD9833] Output disabled (RESET + SLEEP1 + SLEEP12)
[AD9833] Hardware initialized in safe-off state

Boot output was near zero and did not sit near 3.3 V.

Result: PASS

## Test 2 - Connect and synchronize

The PocketLab app connected, enabled response notifications, issued GET_STATE, and synchronized the UI with the firmware.

[RX] GET_STATE
[TX] STATE FREQ=0;AMP=0.00;OFFSET=0.00;WAVE=DC;OUTPUT=OFF

The app and firmware reported the same safe idle state, and connecting did not start the output.

Result: PASS

## Test 3 - Apply valid settings while stopped

A 1 kHz, 0.65 Vpp sine configuration was applied while output remained stopped.

[RX] SET_STATE FREQ=1000;AMP=0.65;OFFSET=0.00;WAVE=SINE
[AD9833] Applied safe-off state; logical WAVE=SINE;FREQ=1000
[TX] OK

The logical settings were stored, but no 1 kHz waveform appeared before RUN.

Result: PASS

## Test 4 - Run output

The app issued OUTPUT ON. The firmware programmed the AD9833 and enabled the waveform.

[RX] OUTPUT ON
[AD9833] Applied FREQ=1000;WAVE=SINE;OUTPUT=ON
[STATE] Output enabled
[TX] OK

A clean 999.97 Hz sine wave measuring 0.59943 Vpp was produced.

Result: PASS

## Test 5 - Stop output

The app issued OUTPUT OFF while the sine wave was active.

[RX] OUTPUT OFF
[AD9833] Output disabled (RESET + SLEEP1 + SLEEP12)
[AD9833] Applied safe-off state; logical WAVE=SINE;FREQ=1000
[TX] OK

The output immediately returned to the millivolt-level safe-off noise floor.

Result: PASS

## Test 6 - Disconnect while output is active

The waveform was enabled and verified before forcing a BLE disconnect.

[BLE] Client disconnected: 51:f5:56:71:3a:f3, reason: 531
[AD9833] SPI write 0x21C0
[AD9833] Output disabled (RESET + SLEEP1 + SLEEP12)

BLE loss automatically disabled the output. The measured signal fell from 0.59943 Vpp to 2.9972 mVpp.

Result: PASS

## Test 7 - Automatic reconnect

The app automatically reconnected, restored notifications, and issued GET_STATE.

[RX] GET_STATE
[TX] STATE FREQ=1000;AMP=0.65;OFFSET=0.00;WAVE=SINE;OUTPUT=OFF

The prior waveform settings were retained, but OUTPUT remained OFF and the signal did not restart.

Result: PASS

## Test 8 - Invalid RUN from DC idle state

After a power cycle, the device was left in DC / 0 Hz / 0 Vpp idle state and RUN was attempted.

[RX] OUTPUT ON
[TX] ERR DC_OUTPUT_NOT_IMPLEMENTED

The invalid request was rejected and the physical output remained near zero.

Result: PASS

# 4. Known-Good Sine Comparison

| Measurement | Previous baseline | Regression test | Difference |
| --- | --- | --- | --- |
| Maximum | 0.641 V | 0.64378 V | +2.78 mV |
| Minimum | 0.041 V | 0.04435 V | +3.35 mV |
| Average | 0.341 V | 0.34371 V | +2.71 mV |
| Peak-to-peak | 0.600 V | 0.59943 V | -0.57 mV |
| Frequency | 999.96 Hz | 999.97 Hz | +0.01 Hz |

The active sine output closely matched the previously measured known-good AD9833 baseline, confirming that the safe-state changes did not degrade normal signal generation.

# 5. Pass-Criteria Evaluation

| Requirement | Result |
| --- | --- |
| Boot output is safe | PASS |
| STOP produces safe-off output | PASS |
| Disconnect forces safe-off output | PASS |
| Reconnect does not restart output | PASS |
| App state matches firmware state | PASS |
| Invalid DC RUN is rejected | PASS |
| Sine output works normally after reconnect configuration is retained | PASS |

# 6. Conclusion

The PocketLab safe-state regression test passed in full. Firmware version 0.3.2 reliably places the AD9833 into a safe-off condition at boot, after STOP, and immediately following a BLE disconnect. Automatic reconnection restores communication and state synchronization without restarting the output. Valid sine-wave operation remains consistent with the known-good baseline, while unsupported DC output requests are correctly rejected.

The only observed software issue was that the mobile app logged the expected DC protocol error as an uncaught promise. This does not affect hardware safety, but the app should later catch the error and present a user-friendly message.