TEST REPORT

PocketLab Amplitude/Offset Interaction Test

Firmware envelope validation versus measured analog output limits

Project: PocketLab BLE-Controlled Function Generator

Test date: July 16, 2026

Test frequency: 1 kHz

Primary waveform: Sine

Overall result: PASS WITH FOLLOW-UP ACTIONS

# Executive Summary

The amplitude/offset interaction test confirmed that the firmware's combined unipolar-output envelope protection behaves as configured. Legal combinations were accepted, out-of-envelope combinations were rejected, and rejected commands preserved the preceding valid waveform. The mobile application also clamped an attempted 4.16 Vpp command to its 4.15 Vpp maximum.

The analog output remained visually sinusoidal without obvious hard clipping across the tested legal combinations. Amplitude accuracy was strong at 2.00 Vpp and 4.15 Vpp. The principal finding was a systematic negative offset error: measured OP484 output mean was typically below the requested offset, with the error increasing to approximately 0.14-0.19 V in demanding medium- and maximum-amplitude cases.

# Test Objective

Verify that the firmware's legal amplitude/offset envelope matches the practical analog limits of the AD5626/X9C103S/OP484 signal path, including combinations near the lower and upper rails and commands immediately inside and outside the configured envelope.

# Configuration and Acceptance Criteria

Configured active-output envelope: 0.20 V minimum to 4.40 V maximum.

Maximum commanded amplitude: 4.15 Vpp.

A legal sine-wave state must satisfy offset - Vpp/2 >= 0.20 V and offset + Vpp/2 <= 4.40 V.

Accepted commands should produce the requested waveform without obvious clipping or rail excursion.

Rejected commands should not alter or momentarily disturb the preceding valid output.

# Measurement Channels

Channel 1: AD5626 offset-control output

Channel 2: OP484 final analog output

# Recorded Test Results

For rejected commands, the displayed measurements belong to the preceding valid state and demonstrate output retention; they are not measurements of the rejected request.

| Test | Request Vpp / Offset | Control result | CH1 Avg / Vpp (V) | CH2 Mean (V) | CH2 Vpp (V) | CH2 Min-Max (V) | Assessment |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 0.50 / 0.46 | Accepted | 0.456 / 0.082 | 0.442 | 0.447 | 0.220-0.667 | Clean; Vpp 10.6% low |
| 2 | 0.50 / 4.14 | Blocked | 0.456 / 0.081 | 0.442 | 0.449 retained | 0.220-0.669 | App sent 4.09 V; firmware OFFSET_OUT_OF_RANGE |
| 3 | 2.00 / 1.25 | Accepted | 1.240 / 0.300 | 1.199 | 2.023 | 0.192-2.215 | Clean; lower peak 8 mV below 0.20 V model |
| 4 | 2.00 / 2.30 | Accepted | 2.283 / 0.301 | 2.206 | 2.016 | 1.195-3.211 | Clean; mean 94 mV low |
| 5 | 2.00 / 3.35 | Accepted | 3.327 / 0.296 | 3.215 | 2.019 | 2.204-4.222 | Clean; mean 136 mV low |
| 6 | 4.15 / 2.30 | Accepted | 2.271 / 0.589 | 2.125 | 4.144 | 0.062-4.207 | Full amplitude; little lower-rail margin |
| 7 | 0.50 / 0.44 | Rejected | 2.270 / 0.589 | 2.114 retained | 4.145 retained | 0.051-4.195 | Correct rejection; prior output retained |
| 8 | 0.50 / 4.16 | Blocked | 2.270 / 0.589 | 2.114 retained | 4.145 retained | 0.050-4.195 | App sent 4.09 V; firmware OFFSET_OUT_OF_RANGE |
| 9 | 2.00 / 1.21 | Accepted | 1.200 / 0.296 | 1.163 | 2.006 | 0.169-2.175 | Clean; measured minimum below model boundary |
| 10 | 2.00 / 1.19 | Rejected | 1.200 / 0.297 | 1.162 retained | 2.005 retained | 0.169-2.174 | Correct rejection; prior output retained |
| 11 | 2.00 / 3.39 | Accepted | 3.367 / 0.293 | 3.251 | 2.018 | 2.240-4.258 | Clean; mean 140 mV low |
| 12 | 2.00 / 3.41 | Rejected | 3.368 / 0.294 | 3.259 retained | 2.013 retained | 2.247-4.260 | Correct rejection; prior output retained |
| 13 | 4.15 / 2.28 | Accepted | 2.252 / 0.589 | 2.107 | 4.145 | 0.044-4.189 | Full amplitude; lower peak near ground |
| 14 | 4.15 / 2.27 | Rejected | 2.250 / 0.589 | 2.096 retained | 4.143 retained | 0.035-4.178 | Correct rejection; prior output retained |
| 15 | 4.15 / 2.32 | Accepted | 2.290 / 0.589 | 2.136 | 4.147 | 0.072-4.219 | Full amplitude; waveform visually smooth |
| 16 | 4.15 / 2.33 | Rejected | 2.289 / 0.589 | 2.132 retained | 4.143 retained | 0.068-4.211 | Correct rejection; prior output retained |
| 17 | 4.14 / 2.30 | App -> 4.15 | 2.270 / 0.589 | 2.113 | 4.147 | 0.049-4.196 | App normalized request to 4.15 Vpp |
| 18 | 4.16 / 2.30 | App -> 4.15 | 2.270 / 0.589 | 2.113 | 4.146 | 0.050-4.196 | App clamped above-maximum request; firmware not exercised |
| R1 | 4.15 / 2.28 | Accepted repeat | 2.250 / 0.590 | 2.098 | 4.147 | 0.034-4.181 | Repeat confirms stability |
| R2 | 4.15 / 2.32 | Accepted repeat | 2.290 / 0.589 | 2.134 | 4.149 | 0.069-4.218 | Repeat confirms stability |

# Analysis

## Firmware Envelope Behavior

All combined-envelope boundary cases that reached the firmware produced the expected outcome. Commands just inside the lower and upper boundaries were accepted, while commands just outside were rejected. The retained-waveform measurements following Tests 7, 10, 12, 14, and 16 show that rejection did not replace the valid operating state.

## Amplitude Accuracy

At a 2.00 Vpp request, measured output ranged from 2.0062 to 2.0228 Vpp across the accepted offset settings, corresponding to approximately +0.3% to +1.1% error. At a 4.15 Vpp request, accepted and repeat measurements ranged from 4.1444 to 4.1490 Vpp, within approximately 6 mV of the request. The 0.50 Vpp low-offset case measured 0.4472 Vpp and therefore showed substantially greater relative error at the low-amplitude end.

## Offset Accuracy and Analog Margin

The measured OP484 mean was consistently lower than the requested offset. Representative errors were -47 mV at 2.00 Vpp/1.21 V offset, -94 mV at 2.00 Vpp/2.30 V offset, -140 mV at 2.00 Vpp/3.39 V offset, and approximately -164 to -184 mV for maximum-amplitude cases.

Consequently, the real analog peaks do not track the firmware's calculated 0.20-4.40 V envelope symmetrically. Maximum-amplitude legal states produced approximately 0.03-4.22 V. Although no obvious hard clipping was visible, the lower peak approached ground closely enough that the current envelope provides little practical lower-rail margin.

## Mobile-App Limit Behavior

An attempted 4.16 Vpp command was forced to 4.15 Vpp on the mobile side before transmission. This is a successful application-layer safety behavior, but it means the firmware's independent rejection of amplitudes above MAX_AMPLITUDE_VPP was not exercised through the normal UI. The intended high-offset low-amplitude requests were also altered to 4.09 V and rejected as OFFSET_OUT_OF_RANGE, revealing a limit or numeric-control mismatch between the app and firmware.

# Conclusions

PASS: Combined amplitude/offset envelope validation operated correctly for all commands that reached the firmware.

PASS: Rejected commands preserved the preceding valid output state.

PASS: No obvious hard clipping or unsafe rail excursion was observed in the accepted cases.

PASS: Medium- and maximum-amplitude accuracy was strong and repeatable.

FOLLOW-UP: Offset calibration does not presently make the requested offset equal the measured OP484 mean.

FOLLOW-UP: The high-offset UI behavior and firmware absolute offset range should be aligned.